## Slide 1: Low Temperature Thermometry on a Helium Evaporation Refrigerator
Tristan Anderson UNH 2020
Advisors: Professor Karl Slifer, Professor Ellie Long
Background
In this project I have designed, installed and calibrated a temperature monitoring system on a helium evaporation refrigerator in Slifer Laboratory at The University of New Hampshire in DeMeritt Hall. This project uses a combination of carbon ceramic and clay thermistors to determine the temperature of eleven locations on the evaporation refrigerator
This thermometry measures the temperature of the Nuclear Physics Group’s (NPG) proton target material and helium refrigerator inside the NPG’s cryostat, Dynamic Nuclear Polarization (DNP) techniques are used by the UNH NPG to create proton targets with high nuclear spin polarizations. Knowing the temperature of the refrigerator and target material is critical to calculating the thermal polarization of a target. The thermal polarization (at thermal equilibrium) of the proton target the NPG is creating can be represented by the hyperbolic tangent function below:
Polarization
Where is the magnetic field, k is the Boltzmann constant, T is the temperature of the material, and is the magnetic moment of the target. The goal of DNP is to maximize the function above by making the numerator large, and the denominator small. Therefore, knowing the temperature of the polarized target material is essential in the NPG’s pursuit of using DNP to create polarized proton targets with high spin polarizations for Jefferson and Fermi National Laboratories.
What is a Thermistor?
A thermistor is a passive electrical resistor whose resistance can be modeled by a function of temperature. A 'normal' thermistor's resistance increases as it gets warmer, allowing it to have a high resolution of temperature at hotter temperature. In Slifer Lab, we use negative temperature coefficient thermistors, which means that as the temperature of the thermistor decreases its resistance increases dramatically, giving an extremely fine temperature resolution below 20 Kelvin. Since Slifer Lab is working with cryogens like Liquid Nitrogen and Liquid Helium at 77 Kelvin and 4.2 Kelvin respectively, it makes sense to choose a negative temperature coefficient thermistor for this system. To the right, is the resistance dependent curve, and pictures of the Allen Bradley (AB), and Carbon Ceramic (CCS) thermistors that we use in the DNP system.
What enables a thermistor to be calibrated, is a functional form that relates resistance to temperature.
The equation above has three unknown variable A, B, and C. These coefficients are solved for by creating three equations with three unknowns. These three equations correspond to the observed resistance and known STABLE temperatures of a particular thermistor at a particular resistance. Slifer lab chooses room temperature, liquid nitrogen, and liquid helium temperatures and resistances to fit the curve above. These three equations are then given to a python program that uses a SciPy library and method called curvefit which solves for the unknowns A, B, and C.
Methods
Planning:
  Material List
  Thermistor Locations
  Wiring paths
Construction and attachment of Thermistors:
  Thermal Anchors
  Streamline wiring
Data Collection:
  LabView collects data from two LakeShore 224 Temperature Monitors and writes raw data to a .lvm file
Calibration:
  Python Script
Procedure of Calibration
Results
Precise, accurate, reliable thermometry throughout the helium refrigerator.
Calibration programs that free up hours of time for other work in the laboratory.
Future Work
Bibliography
Quintell, E. H. "The Low Temperature Characteristics of Carbon-Composition Thermometrs." Review of Scientific Instruments. By J. R. Clement. Vol. 23. New York: American Institute of Physics, 1952. 213-16. Print.
Step 4: That calibration files are uploaded to the Lakeshore 224 temperature monitor
Step 1: Raw resistance data is collected from LabView from an ethernet connection between a computer and the Lakeshore 224 Temperature Monitor
Step 2: The raw resistance data (after collection) is fed into a python program that requires minimal human interaction to clean the raw data
Step 3: Python program uses statistics to select the regions of data to be averaged, then it uses the average of a stable region as a calibration point for the equation below. Scientific python then uses curvefit to solve for the unknowns of the function below.
Create program documentation
Write URA research findings
Proton sample irradiation
~1K was reached in Slifer Lab’s Helium Refrigerator
19-pin connector that mounts to the refrigerator
Size of the thermistors that have been installed
Helium vapor from Helium Dewar
Helium Refrigerator’s lowermost heat exchangers
Position of thermistors on the refrigerator
Helium refrigerator during construction phase